‘How to’ – GeoTribology Project

UMT & NPFLEX

Written by Mia Holt (gy22mh@leeds.ac.uk)

The following methodology is designed for the Bruker Universal Mechanical Tester (UMT) found in the Mechanical Engineering G51a (Dowson Tribology Lab). The document can be used alongside the ‘UMT Standard Operating Procedure’ written by Dr. Joshua Armitage (J.Armitage@leeds.ac.uk).

# 1. Sample Preparation

Sample Material: Carbonate samples - Cararra marble (or limestone)

Tiles of this lithology available to be ordered online (email s.piazolo@leeds.ac.uk). Both materials should be homogenous to reduce effect of extraneous variables (fossils fragments, coarse grains etc.).

## 1.1 Marble/Limestone Plates (Lower Samples)

### 1.2.1 Cutting

Large limestone or marble plates cut into 25 mm x 25 mm squares in Earth and Environment lab (email c.p.woodhouse@leeds.ac.uk) on the small table saw. If the saw is not available contact Dr. Cat Probert (C.M.Probert@leeds.ac.uk) in Civil Engineering and they will organise for the samples to be cut for you.

### 1.2.2 Polishing

Individual 25 mm x 25 mm ‘plates’ are polished on the polisher-grinder machine upstairs in the Thermofluids lab in G.50 in Mechanical Engineering. The samples are ‘polished’ to create grooves on the surface that impose a roughness onto the surface.

- Manually polish to P60, P240 or P1200 on SiC grit paper for 1 minute

- If a ‘smoother’ fault surface roughness is desired, grind to P1200

- Check the sample for tilt and re-polish until the surface is uniform

- Label the sample with an arrow perpendicular to the ridges to maximise roughness

- The direction of the arrow indicates the sliding direction

## 1.2 Marble/Limestone Pins (Upper Samples)

### 1.2.1 Cutting

Large limestone/marble plates of same lithology to lower samples are cut into 5 x 10 cm blocks in Earth and Environment lab (contact Charles Woodhouse - c.p.woodhouse@leeds.ac.uk). If the saw is not available contact Dr. Cat Probert (C.M.Probert@leeds.ac.uk) in Civil Engineering and they will organise for the samples to be cut for you.

Individual pins are cut on a pillar drill with a nominal core drill bit of 10 mm in the Earth and Environment lab (contact Charles Woodhouse c.p.woodhouse@leeds.ac.uk). To cut the individual pins:

- Clamp one block into the vice and submerge in a water-filled plastic container

- Align the block underneath the drill (fitted with a 10 mm drill bit)

- Turn the drill on and lower to the sample surface and apply gentle pressure

- Periodically raise the drill and spray the drill bit with water to prevent clumping

- Each pin takes ~7-10 minutes to cut

- One block can have ~8-10 pins cut from it

### 1.2.2 Polishing

Pins can be polished on the UMT or on the polisher-grinder machine. The UMT is beneficial for inducing a flatter surface, however it can create slight cylindrical geometries to the samples.

#### 1.2.2.1 UMT Polishing Method:

This method uses the rotary drive mounted with SiC paper cut into small circles. The purpose is to create surface roughness and reduce the length of the pins. The pin length from cutting is ~ 10 mm.

- Login: UMT User - ifsumt01

- Open Bruker app on bottom of screen

- Open options file: LIN50G

- Power on UMT

- Load the options file – ‘ROT 50G’

- Install in the rotary drive into the UMT test housing and Mount p60, p240 or p1200

- Ensure that the drive spindle has aligned with the drive’s connector so that the drive sits flat in the housing unit

- Press and rotate the three mounting pins a quarter-turn clockwise to lock them into place

- A TriboID message will appear on the UMT software – ‘ROT Drive Installed’

- Activate motor on UMT

- Open the ‘Grinding sequence script’ in GeoTrib > Mia Holt > Scripts

- Initially, grind all pins using the ‘(P60) calcite, no suspension, grinds to 8mm’ section of the script

  - The pins are now p60, reduced length (~8mm)

- To polish to p240 or p1200 use the ‘(P1200) calcite, no suspension’ or ‘(P240) calcite, no suspension’ section

- Label the pins with an arrow along the side perpendicular to the ridges (same as the lower samples)

  - Ensure the writing on the pins is not on the bottom as this will incorporate graphite into the samples upon sliding

#### 1.2.2.1 Grinder-Polisher Method:

Individual 10 mm pin cores can be polished on the polisher-grinder machine upstairs in the Thermofluids lab in G.50 in Mechanical Engineering. The samples are ‘polished’ to create grooves on the surface that impose a roughness onto the surface. A lower rotation

- Manually polish to P60, P240 or P1200 on SiC grit paper for 1 minute

- If a ‘smoother’ fault surface roughness is desired, grind to P1200

- Check the sample for tilt and re-polish until the surface is uniform

- Label the pins with an arrow along the side perpendicular to the ridges (same as the lower samples)

  - Ensure the writing on the pins is not on the bottom as this will incorporate graphite into the samples upon sliding

## 1.3 Fluid Preparation

Fluid can be made and prepared in the Thermofluids lab in G.50 in Mechanical Engineering. The selected fluid chemistries in these experiments are CaCO3, Na2SO4, MgSO4, BaSO4. Each fluid chemistry has a background of CaCO3 at a supersaturated concentration to enable complete equilibrium. All powders are dissolved in 70 ml of DI water to allow for fluids fill the bath and to be collected before testing.

- Thoroughly rinse a small container to ensure no contamination

- Add 70ml of DI water into a small container

- Add CaCO3 into DI water

- Add MgSO4 or BaSO4 or Na2SO4 into the solution

| 0.4 M | CaCO3 | Na2SO4 | MgSO4 | BaSO4 |
| --- | --- | --- | --- | --- |
| Powder per 70 ml solution (g) | 0.021 | 3.98 | 3.37 | 0.02 |

--------------------------------------------------------------------------------------------------------------

# 2. UMT Set Up

The following instructions are designed for fluids testing on the linear drive for geotribology experiments. Detailed general information about the equipment can be found in Standard Operating Procedure next to the UMT or on Microsoft teams.

- Login: UMT User - ifsumt01

  - If the computer is already on, ensure the adjustments panel and scripts are closed

- Open Bruker app on bottom of screen

- Open options file: LIN50G

- Power on UMT

## Linear Drive & Lower Sample Holder

- Place the linear drive into the UMT (ensure the encoder cables are not squished!)

- Plug the cable into Y socket

- A TriboID message will appear on the UMT software – ‘LIN Drive Installed’

- Attach the adaptor plate to linear drive via countersunk screws #10-32 in three central holes

- Attach the PEEK base plate and stainless-steel heating pad holder via x4 #8-32 UNC screws

- Place the heating pad in the holder and connect the cable to the rear of the large grey heating unit next to the UMT (ensure the unit is OFF)

- Attach the lower sample holder (bath) via x10 #8-32 UNC screws around the edge of the bath

  - Position the cross at the bottom of the bath towards the back of the UMT

- Insert thermocouple no.1 into the small round slot next to the heating pad and plug into the rear of the heating unit (thermocouple 1 slot)

## Load Cell & Upper Sample Holder

- Remove suspension from 50G load cell (if mounted)

  - This is removed to avoid interference with true stick-slip behaviour

- Attach upper sample holder to piezoelectric golden plate of the load cell via x4 #4-40 short screws

  - Ensure the screws are short to prevent damage to the load cell!

  - The arched, strongest section of the sample holder should be at the rear during sliding

- Install the load cell into the UMT housing by sliding it onto the slot below the slider on the carriage

- Fasten it into place using the x4 thumb screws and the lateral handle

- Insert the load cell cable into place

  - A TriboID message will appear on the UMT software – ‘LOAD50g Installed’

- Activate motor on UMT by pressing the motor power button

-------------------------------------------------------------------------------------------------------------

# 3. Running a Test

This stage is ready to be started once the samples are prepared and the UMT is set up and activated.

## 3.1 Before the Test

- Weigh the upper (pins) and lower (plates) samples on weighing scales (3 decimal places)

- Place the upper and lower samples in a supersaturated CaCO3 solution to equilibrate

- Fasten the pH probe into a clamp and plug into the calibrated pH meter

Heating Bath and Fluid Separately

- Plug in thermocouple 2 into the rear of the heating unit and rest on the inside of the empty bath

- Switch on the heating unit and set the temperature to 40 °C

- Decant 70 ml pre-mixed fluid of the desired composition into a glass beaker on a heating and stirring plate

- Place a plate on top to prevent evaporation

- Heat up and stir fluid in beaker to 40 °C on a hot plate (check it is dissolved)

Adding Samples

- Once the bath and fluid are separately heated, fasten the lower sample holder into the bath via x4 #6-32 UNC screws in the square clamps (careful of the heat!)

  - Ensure the direction of the arrow points to front of the UMT

  - Make sure the sample is completely secure - if it is loose it will dislodge during the test

- Fasten the upper sample into the upper sample holder, also ensuring the direction on the arrow points to front of the UMT

  - Ensure this arrow matches the direction of the plate – this is important for analysis post testing

- Pour ~40 ml of fluid into the pre-heated bath ensuring the sample is completely covered

- Fasten the lid of the bath on with x6 #6-32 UNC screws with the thermocouple resting in the thin metal divert near the sample

- Ensure the thermo coupled is rested in the fluid and not touching any metal

  - Position the long slot towards the front of the UMT – this is where the pH probe will go

- Rinse the pH probe with DI water and place into the fluid through the long slot

  - During testing the linear drive will move – so make sure the pH probe has flexibility to move along the slot without colliding with the sample

- Lower the upper sample holder to be almost in contact with the lower sample

  - Adjust the pin to sit towards the rear of the UMT

  - Careful! This process has little room for error - do not dislodge the bath and avoid metal on metal contact

- Take the first pH probe, room humidity and room temperature reading for ‘0 minutes’ when the bath reaches 40 °C

- Open Force (Fx and Fz) graph and adjust as necessary

- Open UMT Test sequence – GeoTrib > Mia Holt > Scripts > V2 Sliding Sequence

- Press red tick to choose test type (ensure only one is ticked)

- Click ‘verify script’

- Start the test via the script and start the timer for measurements

## During the Test

- Decant the cooled remaining fluid from the glass beaker (~15 ml) into a test tube for storage

  - Label with sample name, conditions, and ‘BEFORE’ to ensure samples can be distinguished

  - This is your pre-test fluid

- Record pH, room temperature and humidity:

  - Every minute for the shorter 0.1 mm/s and 1 mm/s tests

  - Every minute until 10 minutes for the longer 10 µm/s tests and then every 30 minutes

## 3.3 After the Test

- Record end time and final room temperature and humidity

- Turn off the heating unit and unscrew the lid of the bath

- Take an image both samples in situ

- Syringe the fluid out of the bath into a test tube labelled with sample name, conditions and ‘AFTER’

- Remove the samples and leave to dry on a heated plate

  - Ensure the gouge is not touched during this process

  - Heat until thoroughly dry so there is no effect of fluids in the final weigh

- Weight the samples when completely dry

- Thoroughly clean the bath with at least 5 rinses and flushes with acetone and DI water until clean

  - This is crucial to remove any contamination between fluid chemistries

- Export data to text file & CSV

- Copy data to one drive

- Close all tabs open on Screen (check hidden on bottom left too)

- Dismantle equipment

- Shut down computer

--------------------------------------------------------------------------------------------------------------

# NPFLEX

The NPFLEX is a non-contact profilometer that performs white light interferometery. It is found in the metrology lab and managed by Ms Camille Hammersley (contact C.Hammersley@leeds.ac.uk). This is used to characterise surface roughness at different resolutions. The settings detailed below have been developed specifically for GeoTribology samples to ensure data points are above 95 % (scientific accepted standard).

## 5.1 Data Acquisition

- Press the green button on the NPFLEX to turn it on

- Turn on the computer monitor

- Load the Vision64 software

- Errors may appear but press okay as this does not affect us

- Click through the notifications that appear

- Place the sample on top of a metal block and secure with playdo

- Change the microscope and magnifier to x20 and x1

- Move the Z axis down to its maximum lowest position to avoid it colliding with the sample after it is moved down

- Unfasten the two handles on the side of the NPFLEX and use the speed dial below to adjust the moving speed

- Use the metal switch down to move the microscope to ~ 0.5 cm above the sample

- Fasten the handles to secure the NPFLEX in place

- Apply the following settings under the ‘Measurement Parameters’ Menu

  - Scan Speed: 1x // lower the higher data no. data points

  - Threshold: 0.5 % // higher sensitivity

  - Backscan: 200 µm // distance above x, adjust depending on prevalence of gouge

  - Length: 200 µm // distance below x, adjust depending on prevalence of gouge

  - Illumination: Green (narrow band) // used for trickier samples

- Select auto-intensity and use the z axis controls to adjust the microscope up and down until the reflectance band shows on screen

  - Use the left most wheel on the control deck to focus this pattern

- For individual scans press ‘Single Acquisition’ in the upper left panel

- For stitched scans go into stitching and change the bounding box to that required then press ‘Measurement’ in the upper left panel

In these samples I do:

  - Length along Scar: 0.3 mm height x 25 mm width

  - Width cross-section of scar: 0.3 width x 25 mm height

## 5.2 Data Analysis and Cleaning:

The following data cleaning instructions are in line with the ISO Standards. This can be done through appsanywhere > vision64, or in the metrology lab on the NPFLEX.

- Apply ‘basic stats’ under 3D analysis and check data points is over 95%

- Apply ‘Terms Removal (F-Operator)

  - Right click > settings > Tilt only // This removes tilt from sample

- Apply ‘Gaussian Regression Filter’

- Right click and all values required will be in here

- You can save as either short or long wavelength

- Apply ‘S Parameters Height’

  - This will tell you the final values after filtering either short or long wavelength

## 5.3 Export Data

To save the Vision64 data file (raw no filters) to your folder, simply click ‘save’ in the top panel as export as an OPDX. To save a specific filtering step/output of data such as s-parameters after filtering, select this step and then right click the map on the screen and click ‘Save Data S-Parameters’.

To export to MatLab you must apply the ‘Data Fill’ and ‘Data Restore’ option after you have filtered your data to be able to run it through the code Josh has written (found in Teams). I recommend using this code to obtain the 3D plots of data but take the roughness values from the non-filled data values under the Gaussian Regression Filter as this is more accurate than using data fill. This code only works with ASC files so make sure to save this accordingly.

Finally, to get data off the computer get the USB locking key from the safe at the end of the lab. Then get a clean USB stick (important!) and transfer your data.